## Slide 1: 3D printed Micropipette

## Slide 2: Cost
Commercial Micropipette
100-1000uL $371.75
10-100ul $384.58
Test Micropipette
BODY+Plunger Printed $124
Nut and bolt $0.30
Spring $0.79
3 mL Syringe $0.15
1mL syringe $0.15
Total with two syringes= $125.39

## Slide 3: Parts
BODY
Plunger
Nut and Bolt
Spring
1mL Syringe

## Slide 4: Assembled

## Slide 5: Design

## Slide 6: Specifications
With the 3mL Syringe Attached
  Dispense up to 1 ml or 1000ul
  Depending on the syringe:
    Measure in increments of 100ul
With 1ml SYRINGE Attached
  Dispense up to .35mL or 350uL
  Measure in increments of 10uL

## Slide 7
Demo:
Dispensing 200uL of water
Scale reads 0.2011g
Percent Error = 0.55%

## Slide 8: Competing Designs
Biropette: customizable, high precision pipette by tbaden on THingiverse

## Slide 9: Our Pipette vs The Competition
Simplicity
  Preparation
  Operation
Adjustable
Accuracy
Durable
Easily Replaceable/Repairable

## Slide 10: 50 Trials to test drift
We set the pipette to about 200 microliters
Dispensed 50 times
Regression Line increased at a small rate
Fix: Replacing the nut with a locking nut
Chart: Trial Number vs Amount Dispensed
 | 0.195 | 0.1955 | 0.196 | 0.1964 | 0.1953 | 0.196 | 0.1954 | 0.1953 | 0.195 | 0.1958 | 0.1955 | 0.196 | 0.1952 | 0.1959 | 0.1944 | 0.1942 | 0.1959 | 0.1942 | 0.1956 | 0.1958 | 0.1958 | 0.195 | 0.1962 | 0.1959 | 0.1958 | 0.196 | 0.1953 | 0.1953 | 0.1955 | 0.195 | 0.1963 | 0.1964 | 0.1958 | 0.1953 | 0.1952 | 0.1959 | 0.1945 | 0.1963 | 0.1955 | 0.1969 | 0.1953 | 0.1963 | 0.1949 | 0.1969 | 0.1958 | 0.1944 | 0.1961 | 0.1946 | 0.196 | 0.1953

## Slide 11: 50 trials testing Drift
Mean | 0.195558
Standard Deviation | 0.000639
Percent Standard Deviation | 0.326872

## Slide 12: Testing
To test the commercial pipettes against ours we performed 15 trials at various amounts. We set our pipette to the measurements using out eyes, and reset the screw before each trial.
To test the 100-1000uL Pipette against ours we set both pipettes to 200uL
  weighed the dispensed liquid in grams
  Theoretical weight : 0.1999g
To test the 10-100uL Pipette against ours we set both pipettes to 50uL and then to 20uL
  weighed the dispensed liquid in grams
    Theoretical weight (50uL) = 0.049999g
    Theoretical weight (20uL) = 0.019999g

## Slide 13: 200uL Commercial Pipette vs Test pipette
Trial # | 100uL to 1000uL Micropipette Weight of 200uL of Water dispensed (g) | Test Pipette (1mL Syringe)
 |  | Weight of 200uL of Water dispensed (g)
1 | 0.2028 | 0.1994
2 | 0.2044 | 0.2001
3 | 0.2014 | 0.1997
4 | 0.2004 | 0.1995
5 | 0.201 | 0.1962

## Slide 14: 50uL Commercial Pipette vs Test pipette
Trial # | 10uL to 100uL Micropipette Weight of 50uL of Water dispensed (g) | Test Pipette (1mL Syringe)
 |  | Weight of 50uL of Water dispensed (g)
1 | 0.0496 | 0.0516
2 | 0.0498 | 0.0514
3 | 0.0498 | 0.0517
4 | 0.05 | 0.0516
5 | 0.0501 | 0.0513

## Slide 15: 20uL Commercial Pipette vs Test pipette
Trial # | 10uL to 100uL Micropipette Weight of 20uL of Water dispensed (g) | Test Pipette (1mL Syringe)
 |  | Weight of 20uL of Water dispensed (g)
1 | 0.0198 | 0.0195
2 | 0.0196 | 0.0196
3 | 0.02 | 0.0196
4 | 0.02 | 0.0195
5 | 0.0199 | 0.0196

## Slide 16: 200uL
Commercial pipette
Average of ten trial : 0.20446
Percent Error: 2.233%
Standard Deviation : 0.00291
Percent Standard Deviation: 1.4215%
Test pipette
Average of 15 trial : 0.19724
Percent Error: -1.377%
Standard Deviation : 0.00485
Percent Standard Deviation: 2.4588%

## Slide 17: 50uL
Commercial pipette
Average of ten trial : 0.0499
Percent Error: -0.197%
Standard Deviation : 0.00014
Percent Standard Deviation: 0.2834%
Test pipette
Average of 15 trial : 0.05315
Percent Error: 6.303%
Standard Deviation : 0.00195
Percent Standard Deviation: 3.6754%

## Slide 18: 20uL
Commercial pipette
Average of ten trial : 0.01983
Percent Error: -0.847%
Standard Deviation : 0.00018
Percent Standard Deviation: 0.8911%
Test pipette
Average of 15 trial : 0.01921
Percent Error: -3.967%
Standard Deviation : 0.00042
Percent Standard Deviation: 2.1948%

## Slide 19
Chart: Percent Error in Amounts Dispensed
| 20uL | 50uL | 200uL
Commercial | 0.00847 | 0.00197 | 0.02233
Test | 0.0251 | 0.0705 | 0.01777

## Slide 20: References
Accuracy and Precision. Ft. Belvoir: Defense Technical Information Center, 1967. Web.
"Discover 3D Models." NIH 3D Print Exchange. N.p., n.d. Web. 16 Sept. 2014.

## Slide 21: Questions