9/03/2014

Measurements for Test Micropipette compared to 100uL to 1000uL and 10uL to 100uL micropipettes

|  |  |  |
| --- | --- | --- |
| Trial # | 100uL to 1000uL Micropipette Weight of 200uL of Water dispensed (g) | Test Pipette (1mL Syringe)  Weight of 200uL of Water dispensed (g) |
| 1 | 0.2028 | 0.1994 |
| 2 | 0.2044 | 0.2001 |
| 3 | 0.2014 | 0.1997 |
| 4 | 0.2004 | 0.1995 |
| 5 | 0.2010 | 0.1962 |
|  |  |  |

|  |  |  |
| --- | --- | --- |
| Trial # | 10uL to 100uL Micropipette Weight of 50uL of Water dispensed (g) | Test Pipette (1mL Syringe)  Weight of 50uL of Water dispensed (g) |
| 1 | 0.0496 | 0.0516 |
| 2 | 0.0498 | 0.0514 |
| 3 | 0.0498 | 0.0517 |
| 4 | 0.0500 | 0.0516 |
| 5 | 0.0501 | 0.0513 |
|  |  |  |
|  |  |  |

|  |  |  |
| --- | --- | --- |
| Trial # | 10uL to 100uL Micropipette Weight of 20uL of Water dispensed (g) | Test Pipette (1mL Syringe)  Weight of 20uL of Water dispensed (g) |
| 1 | 0.0198 | 0.0195 |
| 2 | 0.0196 | 0.0196 |
| 3 | 0.0200 | 0.0196 |
| 4 | 0.0200 | 0.0195 |
| 5 | 0.0199 | 0.0196 |
|  |  |  |
|  |  |  |